**POCT Link Trainer Guide – SureStep pregnancy test training**

This document summarises the key points which should be covered by link trainers during training of **pregnancy testing** using the **SureStep pregnancy test.**

|  |  |
| --- | --- |
| **Purpose** SureStep™ Urine Drug Test Cassette | Abbott Toxicology | * Human chorionic gonadotrophin (hCG) is a glycoprotein hormone, which is produced by the developing placenta shortly after conception. * The SureStep test is a qualitative immunoassay, which detects the intact hCG molecule in urine. |
| Urine | Hull University Teaching Hospitals NHS Trust**Sample requirements** | * The urine sample must be collected into a clean dry plastic or glass container without preservatives. Label with the patient’s details. * Specimens collected at any time of day may be used, but the first urine sample of the day usually contains the highest concentration of hCG. * If the test is not to be done immediately the sample should be labelled with the patient’s full name, DOB and NHS number and stored at 2-8˚C for up to 48 hours. |
| **EQA** | * EQA samples should be tested on receipt, and the results returned to POCT on the form provided ASAP. |
| Image result for WEAR GLOVES**Patient / sample preparation** | * Gain positive patient identification and (when possible) consent. * Wear the appropriate PPE. * Collect a fresh urine sample in a clean container labelled with patient details. |
| **Analysis** | * Check the test expiry date. * Fill in the audit sheet. * Ensure the patient sample is at room temperature. * Remove the pregnancy test strip and dropper from the foil wrapper and place on a flat surface. Label the strip with the patient’s name/NHS number. * Using the dropper provided dispense three full drops of urine into the round sample well. * After 3-4 minutes, read the result. **DO NOT** interpret the results after the 4-minute appropriate read time has passed. |
| **Results** | * Pregnancy (hCG) Rapid Test Cassette(Urine) JusChek Women Health Rapid Test  Kuala Lumpur (KL), Malaysia, Selangor Supplier, Suppliers, Supply, Supplies  | Setia Scientific SolutionTo interpret the results please ensure that you are in good light. * All results (including invalid results) must be recorded in the patients’ notes and on the audit sheet. Failure to do so will result in the test kits being withdrawn. |
| **After testing** | * Dispose of test strips following waste management guidance. |
| **Stock** | * Tests are supplied by POCT when a filled audit sheet is returned and tests are requested. |
| Make a Phone Call**Analyser issues** | * Contact POCT: * York (772)5890 * Scarborough (771)2659 * Bridlington (771) 3321 |

**\* DO NOT SHARE YOUR OPERATOR ID WITH ANYONE \***

**\* IT IS A DISCIPLINARY OFFENCE \***